## Slide 1: An Introduction to the Basics of PET Imaging
N. Jon Shah
Institute of Neuroscience and Medicine – 4
Research Centre Juelich
52425 Juelich
GERMANY

## Slide 2: Outline
Introduction
    PET Basics
    Example Measurement
Overview Processing Workflow
Required Data Corrections in Detail
    Attenuation
    Randoms
    Normalisation
    Compton Scattering
    Decay & Deadtime Correction
    Calibration

## Slide 3: Mental activity needs energy.
Substrates providing energy (glucose and oxygen) must be transported to the brain.
This transport can be observed with PET.
Example of
Functional Imaging Using PET

## Slide 4: H2O
H216O
Water is distributed within the body corresponding to blood flow

## Slide 5: H2O
H216O
H215O
ß+
e-
ß+
Using “radiating“ water blood flow becomes visible
511 keV
511 keV

## Slide 6: HR+ & Photonenpaare

## Slide 7
Dt < 12 ns : coincidence o.k.
BGO
LSO
NaI
GSO

## Slide 8: The ring detector records pairs of annihilation photons by looking into the object from different views
Projections of the
point sources
Two point sources
emitting pairs of annihilation photons

## Slide 9: Projections seen from around the head - activity used: 18F-fluoro-deoxy-glucose

## Slide 10
Image Reconstruction: Backprojection of the measured projection data and superposition
Projections of point sources
Image of point sources

## Slide 11: Rekonstruktionen
Pi = Sij Aj
i = 1 , ... , 288 x 144
j = 1 , ... , 128 x 128 I x J = 679.477.248
Algebraic Reconstruction by
Iterative Solving of Matrix System
Analytic Reconstruction by
Filtered Backprojection

## Slide 12: Metabolic Functions Visualised
HO
18
F
H
O
H
O
CH OH
2
O
-
Metabolic Functions Visualised by “Natural“ Tracer Molecules Labelled
with Positron Emitters
H215O
Perfusion
Glucose Consumption
15O2
Oxygen Consumption
H C
3
HO
O
S
NH
2
11
C
Transport
of Amino Acids

## Slide 13: Examples of reconstructed images
H215O:
Cerebral blood flow
11C-SDZ GLC756:
dopamine
receptors
18F-Altanserin:
serotonin
receptors

## Slide 14: Visualisation of Radioactivity
18F-Altanserin
⇨ serotonin receptors
counts/s
1.50 E06
Visualisation of Radioactivity
1.29 E06
1.07 E06
8.57 E05
6.43 E05
4.29 E05
2.14 E05

## Slide 15: Quantitation of Radioactivity Concentration
16
12
10
8
4
kBq/ml
6
14
Measurement of 18F-altanserin uptake
Quantification of Radioactivity Concentration

## Slide 16
=
mTissue
mSyringe
cSyringe
cPET
From Radioactivity Concentration (kBq/ml)
to Mass Concentration (mmol/ml)

## Slide 17: Quantitation of Mass Concentration
86
66
54
43
22
fmol/ml
33
75
Quantification of Mass Concentration
Measurement of altanserin concentration

## Slide 18: Simultaneous PET-MR patient measurements
18Flour-Deoxy-Glucose-PET
simultaneous MR (T1 MPRAGE)
metabolic
information
anatomical
information

## Slide 19: First Principle: Positron-Emission-Tomographie (PET)
detector
combination
( D1 , D2 )
line of response
(R, q)
D1
D2
g
g
18F
b+ decay process of metabolic tracer:
e.g.: 18F --> 18O + e+ + ne
e+ e- --> 2 g

## Slide 20: 3D data acquisition
144 crystals / block
6 blocks / modul
32 moduls
27648
detector crystals
1 HEAD
19 OPPOSITE HEADS
227 million
detector combinations
(Lines-Of-Response)

## Slide 21: BrainPET Data Formats (I)
List Mode (LM) Data:
list of detected coincidences
in chronological order
1.
2.
3.
33
54
41
7
50
1
2. event: ( 7 ,41)
3. event: (33,54)
1. event: ( 1 ,50)
store pairs of
detector numbers
Acquisition
Time
additional tags are
periodically inserted:
time tags
count rate tags
external trigger tags

## Slide 22: BrainPET Data Formats (II)
Line-of-Response (LOR) Data:
sum of detected coincidences
for any detector pair
( 0, 50) --> #0
( 0, 51) --> #0
( 0, 52) --> #2
......
( 1, 50) --> #1
.....
.....
(33, 54) --> #2
.....
.....
230 million
detector pairs
List Mode
Sorter
List Mode (LM) Data
select data
of specific
period of time
(time frame)
t1
t2
event ( 1, 50)
event ( 7, 41)
event (33, 54)
event (29, 52)
event (17, 20)
event ( 1, 50)
x
x
x
event (33, 54)
x
x

## Slide 23: BrainPET Data Formats (III)
LOR Data
Sinogram Data:
geometrical sorting
of LOR data into projections
33
54
{ D1, D2 }  { R , q }
+2
Sinogram
Sorter
( 0, 50) --> #0
( 0, 51) --> #0
( 0, 52) --> #2
......
( 1, 50) --> #1
.....
.....
(33, 54) --> #2
.....
.....

## Slide 24: BrainPET in 3T-Tim-Trio MR

## Slide 25: BrainPET insert
scanner with 32 heads
38 cm
single detector head
copper shielded against MR influence
6 PET detector blocks / head
20 cm
axial
12 x 12 crystals / block
2.5 x 2.5 x 20 mm3 LSO crystal
3 x 3 APD / block
20 cm
axial
electronics

## Slide 26
Design and Structure of the BrainPET
unrolled detector
72
detector
rings
12x12 crystals, 3x3 APDs

## Slide 27
Measurement Effect: Deadtime
PET countrate:
deadtime effect:
deviation from global correction factor:
0
100
200
300
CFD-singles/s
400
1
1.1
1.2
1.3
1.4
ideal curve
measured curve

## Slide 28
Measurement Effect: Pile-up
APD
C
D
A
B
Anger principle
at low countrate:
at high countrate:
ratio (crystal level):

## Slide 29: 7T BrainPET Insert
Christoph Lerche and Debora Niekämper + FZJ, RWTH, Monash Teams
Thanks to Siemens for providing the crystals!

## Slide 30
PET Scanners – State of the art
BrainPET
Spatial resolution:
  centre ≈ 2.5 - 3 mm
  edge ≈ 5 mm
Axial field-of-view 19.2 cm
Sensitivity ≈ 7%
Combined with 3T MRI
Clinical standard
Spatial resolution:
  centre ≈ 4 mm
  edge ≈ 6 mm
Axial field-of-view 15-25 cm
Sensitivity ≈ 2%
PET-CT Siemens Biograph Vision
Siemens MAGNETOM Trio MRI and the BrainPET insert in Jülich
≈20cm

## Slide 31
Improved sensitivity (>12%)
Extended imaging region axial field-of-view ≈ 25cm
Improved spatial resolution (1.5–2 mm)
Improvement of time resolution (by >90%)
Combined with 7T MRI
7T BrainPET Insert
BrainPET 7T gantry with mounting / transport cart Status August 2022
5 mm3
40 mm3

## Slide 32
Gantry
Check of the integration of modules and assembly
Bar for precise axial and transversal positioning of the insert
Connectors for alarm ball and head phones
Mounting allowing rotation

## Slide 33
Modules (RWTH Aachen)
Carbon fibre housing
Hose for liquid cooling
Feed trough for optical links and air flushing
Mounting and transport box

## Slide 34
Modules (RWTH Aachen)
RF Seal
Cooling liquid distribution
PET Module electronics

## Slide 35: Original J-pole antenna array
Choi et al., IEEE Trans. Med. Imaging, 2022
Choi , Hong, Felder, Shah, German patent filed in 2020
Notes: Multi-channel benefits and decoupling problem …

## Slide 36
PET listmode 0-60 min: Dynamic FDG-PET
T1 / T2/ UTE
ASL
fMRI restingstate
Angio
Hybrid MR-PET
dynamic FDG-PET with MRI
DTI
Notes: FT4799F

## Slide 37: Thank you for your attention Questions?
depth
I
II/III
IV
V
VI
CSF
WM