## Slide 1: PET and PET/MR in Neuroscience
Kutaisi, Georgia, September 12, 2024 I Christoph LErche
Quantitative Hybrid MR/PET Imaging
Notes: I would like to thank everyone who has made it here today and who is interested in my presentation

## Slide 2: Positron Emission tomography in a Nutshell
Operating Principle of Positron Emission Tomography (PET)
[18F]FDG (Flourodeoxyglucose)
Tissue atom
Tracer Principle:
concentration of contrast agent very low
⇒ does not alter biological system
(Nobel Prize Chemistry, 1943, George De Hevesy)
≈ 500･106
Coincidences
LOR histograms
or
Sinograms
Image reconstruction (ML-EM)
(Maximum Likelihood – Expectation Maximisation)
System matrix
(probability)
Coincidence counts in LOR i
Image voxel value j
(decays)
Shepp and Vardi. IEEE transactions on medical imaging 1.2 (1982): 113-122.
Coincidence Unit
Line of Response
(LOR)
Notes: Mention tracer principle here

## Slide 3: Simplified Receptor Occupancy Theory
Competition between (reversible) tracer and endogenous neurotransmitter release
Bound activity concentration
Receptors
Neurotransmitter
Radioligand
(tracer)
Synaptic gap
Postsynaptic neuron
Time
Base-line (steady state)
Receptors
Neurotransmitter
Radioligand
(tracer)
Synaptic gap
Postsynaptic neuron
Endogenous neurotransmitter release (cognitive task)
⇒ Displacement of reversible tracer
Receptors
Neurotransmitter
Synaptic gap
Postsynaptic neuron
20 nm
100×100 nm2
total number of synapses in the brain: ~150×1012
Notes: Average number of neurons in the brain = 86 billion
Average neuron has about 1000 synapses, total number of synapses ~150×1012 trillion
Typical size of synaptic cleft : 100 ×100 × 20 nm3

## Slide 4: PET Application – Neurology/Neuroscience
Imaging neurotransmitters, neuroreceptors, metabolites, and neuroreceptor occupation
[18F]FDG (glucose)
[11C]FMZ (GABAA)
[18F]Fallypride (D1/D2)
[11C]ABP688 (mGluR5)
Kim, Song E., et al. Brain Structure and Function 225 (2020): 1805-1816.
Glutamate
most abundant excitatory neurotransmitter in the human brain
GABA
most abundant inhibitory neurotransmitter in the human brain
Dopamine
relevant in addiction, reward, Parkinson's disease and others
Glucose
Brain energy metabolism
GABA: g-Aminobutyric acid
mGluR5: Metabotropic glutamate receptor 5

## Slide 5: Depression – Example for A multifactorial Disease
Dean, Jason, and Matcheri Keshavan. "The neurobiology of depression: An integrated view." Asian journal of psychiatry 27 (2017): 101-111.
Aetiology
Genetics
Epigenetics
Stress
Trauma
etc.
Neurotransmitters
5-HT (Serotonin)
DT (Dopamine)
NA (Noradrenaline)
Glu (Glutamate)
GABA (Gamma Aminobutyric Acid)
Inflammation
Altered connectivity
(prefrontal cortex – limbic)
Reduced neuroplasticity
Hypothalamic-Pituitary-Adrenal axis hyperactivity
Pathophysiology
Möhler, Hanns. "The GABA system in anxiety and depression and its therapeutic potential." Neuropharmacology 62.1 (2012): 42-53.
Notes: suicides Germany approximately 9000, worldwide: 700.000
The economic burden of MDD (major depressive dissorder) in the United States in the year 2000 was estimated to be 83.1 billion dollars, with 26.1 billion dollars (31%) directly related to medical costs, 5.4 billion dollars (7%) related to suicide, and 51.5 billion dollars (62%) due to workplace costs including absenteeism and reduced productivity
The ability of neurons to modify the strength of existing synapses, as well as form new synaptic connections, is called neuroplasticity. Defined in this way, neuroplasticity includes changes in strength of mature synaptic connections, as well as the formation and elimination of synapses in adult and developing brains.

## Slide 6: Combining Imaging of the Parts of Cerebral Communication
Functional MRI, PET and EEG
Domain of fMRI (BOLD Effect)
Image adapted from https://www.cognitivefxusa.com/blog/neurovascular-coupling-the-secret-to-recovering-from-a-concussion
Neuron
Astrocyte
Blood vessel
Neurovascular coupling
Hemodynamic response
Systemic level
Analysis of centers of cerebral data processing
Domain of PET
(Receptor density/occupancy)
Neurotransmitter
Neurotransmitter
transporter
Receptor
Synaptic vesicle
Voltage-
gated
channel
post synapse
Axon terminal
Synaptic cleft
Synapse
Molecular level
Analysis of chemical neurotransmission
Electrode
Cortex
Electrical field
Cranium
Image adapted from S. Beniczky, et al. Epileptic Disorders
22.6 (2020): 697-715 and L. Peicha, “Movement Decoding from EEG:
Target or Direction?”, MSc. Thesis, Institute of Neural Engineering, Graz, (2017)
Domain of EEG
(Evoked potentials)
Electrical level
Analysis of temporal signatures of neurotransmission
Notes: PET: Can be used to measured the receptor availability under different conditions, for instance before and after pharmacologic or cognitive challenges
fMRI: reflects the oxygen extraction from the blood and is therefore expected to be linked to the activity of neurons by oxygen consumption
EEG: reflects voltage changes resulting from synchronous firing of groups of neurons

## Slide 7: Advanced MRI Methods
functional MRI, Spectroscopy, and X-Nuclei Imaging
Madelin, et al. JMRI, 2013
Example: Sodium Imaging
Nuclei 23Na, 31P, 19F, …
With courtesy of E.Farherr, INM-4, Forschungszentrum Juelich GmbH
Energy metabolism
Glucose
Lactate
Alanine
Phosphocreatine
Creatine
Neurotransmitters
Glutamate
Glutamine
GABA
Aspartate
Glycine
Serine
N-acetylaspartate-glutamate
NeuroImage 168 (2018) 181–198
MRS at 7T
Finger tapping study @ 7T
Image courtesy of P. País Roldan, INM-4, Forschungszentrum Juelich GmbH
fMRI at 7T
Others:
Structural imaging with excellent spatial resolution and soft tissue contrast
Vessel imaging
Perfusion weighted imaging
Diffusion weighted imaging (tractography)

## Slide 8: Typical Multimodal MR/PET/EEG Study Protocol
[11C]ABP688 PET + fMRI (Resting state vs. task) + EEG (Event related potentials) + MRS
C. Régio Brambilla, et al. "mGluR5 binding changes during a mismatch negativity task in a multimodal protocol with [11C] ABP688 PET/MR-EEG." Translational Psychiatry 12.1 (2022): 1-10.
L. Orth, “Multimodal Metabolic Imaging Using Single Voxel MRS, DTI and [11C]ABP688 PET in Schizophrenic Patients”, Master Thesis, Research Centre Jülich GmbH, Germany
Del Guerra, Alberto, et al. "TRIMAGE: A dedicated trimodality (PET/MR/EEG) imaging tool for schizophrenia." European psychiatry 50 (2018): 7-20.
Challenge
(tone)
standard tone
deviant tone
Amplitude [µV]
Amplitude [µV]
time [ms]
healthy control
schizophrenic patient

## Slide 9: Spatial Resolution And Sensitivity for Brain PET
13,5 cm
15,5 cm
Image courtesy of Markus Axer, Institute of Neuroscience and Medicine 1,
Forschungszentrum Juelich GmbH
Minimize Partial Volume Effect:
2 FWHM (spatial resolution) ≲ imaged structure
Autoradiography
Zilles et al., Nature Reviews Neuroscience 11.2 (2010): 139.
Receptors mainly located in cortex (grey matter)
Cholinergic muscarinic M2 receptors in human brain (coronal sections)
Requirement for new Development

## Slide 10: Simultaneous MR-PET-EEG Imaging – 2nd Generation
A high performance Ultra High-Field MR compatible PET insert for neuroscience
Target performance parameter | 25 cm
Axial FOV | 25 cm ✓
Sensitivity (point source at centre) | ✓
Spatial resolution (entire FOV) | ✓ (1.6 mm)
Timing resolution | (✓ )
MR compatibility | 3T, 7T & 9.4T, several nuclei (✓ )
Target performance parameter | 25 cm
Axial FOV | 25 cm ✓
Sensitivity (point source at centre) | 
Spatial resolution (entire FOV) | 
Timing resolution | 
MR compatibility | 3T, 7T & 9.4T, several nuclei (✓ )
Siemens Magnetom Terra 7T
High resolution detectors
LSO, 24 mm thickness 3 layers,
2 mm pitch, digital SiPMs
PET Transparent RF Coil
8 channel dipole TX/RX array, PET transparent
Long axial FOV gantry
Liquid cooling, suited for neuroscience
UHF MR compatible PET Modules
Access to raw data and singles
Helmholtz Validation Fund HVF 0051
Innovation Go “BrainPET 7T”

## Slide 11: Brain Phantom Images
No normalization, no random correction, no scatter correction (but with attenuation correction)
Iida, Hidehiro, et al. Annals of nuclear medicine 27 (2013): 25-36.
3D brain phantom with bone & grey matter compartments
Head holder
Activity: ≈ 45 MBq (avg. act. conc. ≈ 12kBq/ml, 18F), acquisition time: ≈ 270 sec.
Reconstruction with PRESTO*, (ML-EM, 70 iterations, voxel size 0.652 × 1 mm3)
*J. Scheins, et al., IEEE transactions on medical imaging 30.3 (2011): 879-892.

## Slide 12: First MR images with running BrainPET 7T
Acquired just last week!
Turbo-flash sequence with head phantom
EPI sequence with head phantom
Gradient echo sequence with fruits

## Slide 13: Imaging Synaptic Density
Isoform Synaptic Vesicle Glycoprotein 2A (SV2A) Provides Access to Plasticity
Rossi, Rachele, et al. "Synaptic vesicle glycoprotein 2A: features and functions." Frontiers in neuroscience 16 (2022): 864514.
Prefusion during vesicular exocytosis
presynaptic synapse terminal
vesicle
vesicle
exocytosis
docking
Modulation of neurotransmitter release
Interaction with laminin
Interaction with other SV proteins, e.g. calcium sensor
Roles of SV2A
Synaptic density
Vesicle density
SV2A protein density
Interpretation of reduced binding
Naganawa, Mika, et al. "First-in-human evaluation of 18F-SynVesT-1, a radioligand for PET imaging of synaptic vesicle glycoprotein 2A." Journal of Nuclear Medicine 62.4 (2021): 561-567.
Vt measurements with PET
pre-fusion
maturation
Notes: Summary of the main putative roles of the isoform synaptic vesicle glycoprotein 2A (SV2A). SV2A participates in a pre-fusion maturation step during vesicular exocytosis; SV2A is involved in neurotransmission via multiple mechanisms, e.g., by modulating the coupling between calcium entry and neurotransmitter release or by regulating machinery involved in neurotransmitter release; SV2A interacts with extracellular matrix components, e.g., laminin; SV2A interacts with other synaptic vesicle proteins, such as the calcium-sensor synaptotagmin. Other SV2A possible functions are reported, for example the transport of galactose, mediation of vesicular entry of neurotoxins, interaction with adenine nucleotides and mitochondrial localization and role as fusion or fission factor.
What does “altered SV2A density” represent? Three possibilities: (A) Altered synaptic vesicle glycoprotein 2A (SV2A) tracer binding could be associated with changes in “synaptic density” (exemplified as a change in the number of dendritic spines). (B) Altered SV2A tracer binding could be associated with changes in the number of vesicles in the pre-synaptic terminal, according to the functional state of the synapse, and so with functional properties of the synapses rather than their density. (C) Altered SV2A tracer binding could be associated with changes in the number of SV2A copies in the vesicles. The number of SV2A copies reached by the tracers and thus labeled may also vary. The alternative scenarios (B) and (C) do not necessarily implicate a change in the overall synaptic density.

## Slide 14: Depression – Example for A multifactorial Disease
Dean, Jason, and Matcheri Keshavan. "The neurobiology of depression: An integrated view." Asian journal of psychiatry 27 (2017): 101-111.
Aetiology
Genetics
Epigenetics
Stress
Trauma
etc.
Neurotransmitters
5-HT (Serotonin)
DT (Dopamine)
NA (Noradrenaline)
Glu (Glutamate)
GABA (Gamma Aminobutyric Acid)
Inflammation
Altered connectivity
(prefrontal cortex – limbic)
Reduced neuroplasticity
Hypothalamic-Pituitary-Adrenal axis hyperactivity
Pathophysiology
Serotonin:
metabolites are reduced in patients diagnosed with depression
Dopamine:
antidepressant agents increase dopamine levels in the brain
Glutamate:
ketamine (receptor antagonist) acts as fast antidepressant
Noradrenaline:
medications that inhibit NA reuptake are effective antidepressants
GABA:
mood amenable to drugs targeting at GABA receptors
Möhler, Hanns. "The GABA system in anxiety and depression and its therapeutic potential." Neuropharmacology 62.1 (2012): 42-53.
Notes: suicides Germany approximately 9000, worldwide: 700.000
The economic burden of MDD (major depressive dissorder) in the United States in the year 2000 was estimated to be 83.1 billion dollars, with 26.1 billion dollars (31%) directly related to medical costs, 5.4 billion dollars (7%) related to suicide, and 51.5 billion dollars (62%) due to workplace costs including absenteeism and reduced productivity

## Slide 15: Dual Tracer PET Imaging (Simultaneously)
Prompt gammas allow for source discrimination of the PET tracer!
b+⇒ 2 g (511 keV)
g (specific energy)
PET Nuclide
emitter
More details in the next talk!
Image with courtesy from D.Niekämper, INM-4, FZJ
Pfaehler, E. (2024). ML-EM based Dual tracer PET image reconstruction with inclusion of prompt gamma attenuation (1.0). Zenodo. https://doi.org/10.5281/zenodo.13382960 (preprint)
Source distribution
(simulation)
Quantitative reconstruction algorithm (including attenuation):
Simulation mimics head including bone
1 iteration
50 iterations

## Slide 16: Layer-specific detection of brain function at 7T
Six cortical layers, each of which has its particular function
Yun, Seong Dae, et al. "Mapping of whole‐cerebrum resting‐state networks using ultra‐high resolution acquisition protocols." Human Brain Mapping (2022).
RH
layers ROI 8
layers
ROI 10
correlation
Connectivity
Super
(near CSF)
Deep
(near WM)
Cortical depth
Default Mode
Visual
Sensory-Motor
Exec. Control
Images with courtesy of P. País Roldan and Yun, Seong Dae

## Slide 17: Receptor Internalization
PET and fMRI time courses contain complementary information about receptor adaptation mechanisms
Christin Y. S ander, Simultaneous PET/fMRI for Imaging Neuroreceptor Dynamics, PhD Thesis MIT 2014
functional hemodynamic response DH DH = f (receptor occupancy)
functional hemodynamic response can be measured with fMRI, e.g. CBV vs. time
fMRI statistical analysis
PET kinetic modelling results
Sander, Christin Y., et al. "Imaging agonist-induced D2/D3 receptor desensitization and internalization in vivo with PET/fMRI." Neuropsychopharmacology 41.5 (2016): 1427-1436.
Blocking study in non-human primates
Notes: Schematic illustrating a compartmental model that describes receptor desensitization and internalization at dopaminergic synapses. The total number of receptors (Bmax) is composed of available receptors at the postsynaptic membrane, those bound by an injected agonist, those bound by endogenous dopamine and desensitized/internalized receptors. Occupied receptors are in exchange with free ligand in the synaptic space. Receptors that are occupied by an agonist trigger desensitization and internalization. As externalization mechanisms are known to be very slow, we assume that kext=0 for the duration of time courses we modeled. The parameters that determine the PET and fMRI signal changes are highlighted in bold. This shows that PET and fMRI time courses contain complementary information about receptor adaptation mechanisms. concentr., concentration; DA, dopamine; fMRI, functional magnetic resonance imaging; PET, positron emission tomography.
Time courses of CBV (dotted line) and occupancy (solid line) resulting from exposure to three different pharmacological challenges in animal M2: (a) the high-affinity agonist quinpirole elicits a very short CBV response, whereas PET occupancy stays elevated for the duration of the experiment. (b) The agonist ropinirole has a lower affinity compared with quinpirole and results in a slightly longer, though still short, CBV response, whereas occupancy peaks at 17.8 min and then starts to decrease. Compared with quinpirole, ropinirole displays a larger CBVpeak signal at lower occupancy. (c) The antagonist prochlorperazine shows that CBV and occupancy time courses are matched, and demonstrates that CBV can stay elevated for longer durations in time. The antagonist showed a reversed CBV sign and the largest magnitude compared with the agonists. Overall, the discrepancy in time between CBV and occupancy and diminished CBV magnitude for the agonists suggests that RDI affects both PET and fMRI, and can vary with drug affinity and potency. CBV, cerebral blood volume; fMRI, functional magnetic resonance imaging; GLM, general linear model; PET, positron emission tomography; RDI, receptor desensitization and internalization
According to the classical occupancy theory, receptor occupancy by agonist causes a downstream functional response, which can be observed with PET/fMRI. High-affinity agonists are known to cause receptor internalization. The functional hemodynamic response (measured as changes in cerebral blood volume (CBV) in this study) to a D2 agonist can be expressed in terms of changes in receptor occupancies of the agonist (ligand L) and endogenous DA, which is displaced according to the law of mass action. The functional CBV signal can be modeled through a direct coupling relationship with occupancy
Parametric maps showing the results from three injection doses of the agonist quinpirole. All maps show imaging data from two animals, analyzed with a mixed-effects model. Upper row: dynamic-binding potential maps (DBPND peak) show that specific binding of the radiotracer [11C]raclopride in caudate and putamen decreases with increasing quinpirole. Lower row: voxelwise maps showing %CBVpeak changes, windowed by a P-value map with Po0.03. As quinpirole dose increases, the negative CBV signal that is specific to caudate and putamen increases in magnitude. CBV, cerebral blood volume; DBP, dynamicbinding potential.

## Slide 18: Take Home Messages
PET Provides Unique Access to Neurotransmission (Chemical Component of Processing)
MR Provides (indirect) Unique Access to Neuron Activity
MR/PET Together allow Simultaneous Multiparametric Functional Imaging
Natural Information Processing in the Brain is Inherently Multiparametric
Most Dysfunctions and Diseases of the Brain are Inherently Multifactorial

## Slide 19: Acknowledgements & Collaborations
PET Group
Prof. H. Herzog
Dr. C. Régio Brambilla
Dr. C. Li
Dr. J. Scheins
Dr. E. Pfähler
Dr. A. Issa
Dr. Qi Liu
H. Alrukh
N. Beriashvili
D. Niekämper
Multimodal Imaging Group
Prof. I. Neuner
Dr. S. Ramikran
Dr. R. Rajkumar
MR Group
Prof. N. Jon Shah
Dr. P. País Roldan
Dr. S.D. Yun
Dr. C.H. Choi
Dr. J. Felder
Dr. J. Mauler
Dr. E. Farrher
S. Schwan
Partners
Prof. G. Natour
Prof. v. Waasen
Prof. V. Schulz
Prof. G. Egan
Dr. D. Schug
Dr. B. Weissler
Dr. Z. Chen
Dr. P. Gebhardt
Dr. F. Müller
H. Radermacher
J. Colliene
D. Grunwald
M. Lennartz
D. Arutinov
R. Heil
W. Silex

## Slide 20
Questions?
INM-4
Medical Imaging Physics
c.lerche@fz-juelich.de

## Slide 21: Artificial Intelligence Is not The new 42
Wikipedia: Science is a strict systematic discipline that builds and organises knowledge in the form of testable hypotheses and predictions about the world.
Science is full of abstractions. A line plot is an abstraction. A false-colour image is an abstraction. Even scientific notation like 6.0 × 1023 is an abstraction. Abstractions like these are central to science, invoked all the time, and easy to understand. (The Goldilocks level of abstraction in science, Ken Hughes)
Britannica: Abstraction, is the cognitive process of isolating, or “abstracting,” a common feature or relationship observed in a number of things, or the product of such a process.
Sciencecouncil.org: A scientist is someone who systematically gathers and uses research and evidence, to make hypotheses and test them, to gain and share understanding and knowledge.
Example: Consider an Exponential function (which is an abstract model for Activity of some radioactive source after some time)
4 layers, 7600 parameters

## Slide 22: Neuroimaging an Artificial Intelligence
Need for growing set of image data sets!
Further obstacles to overcome:
bias is training set: spectrum bias (well defined cohort), image device or imaging condition (e.g. du to legal limitation as radiation protection), labelling errors, lack of understanding
evaluation metrics: normally performance on benchmark data set, no superiority testing as required in clinical trials
robustness to out-of-distribution samples
Availability of large labelled datasets enabled solving difficult ML problems as image recognition with datasets of millions of images ⬄medical datasets are much smaller, normally of thousands or hundreds or even less (remember the legal limitations!)

## Slide 23: Some Articles of Interest
https://spectrum.ieee.org/stop-calling-everything-ai-machinelearning-pioneer-says
https://www.inc.com/jason-aten/stop-calling-everything-ai-its-just-computers-doing-computer-stuff.html

## Slide 24: Clinical PET Application – Neuro-oncology
Diagnosis, staging, and therapy planning of (cerebral) tumors with MRI, [18F]FET, and [18F]FDG
 | MRI | MRI + [18F]FET PET
Sensitivity | 96 % | 93 %
Specificity | 53 % | 94 %
[18F]FDG | [18F]FET
35% | 86%
[18F]FDG
PET
[18F]FET
PET
Comparison between FDG and FET
19 high grade glioma (HGG)
24 low grade glioma (LGG)
Pauleit et al., Nucl. Med. Biol. 2009; Vol. 36(7), pp: 779-87
O-(2-[18F]-Fluoroethyl)-L-Tyrosine (FET)
[18F]-2-Fluor-2-desoxy-D-Glucose (FDG)
Comparison of MRI and [18F]FET PET
[18F]FET PET
T1-MRI
T2-MRI
Pauleit, Dirk, et al. Brain 128.3 (2005): 678-687.
Increased specificity with PET
cases with increased uptake in tumors:

## Slide 25: Noise Equivalent Count Rate
Adapted NEMA: Phantom diameter 15 cm, length 40 cm
 | Phantom | aFOV | CTR | NECRPeak | SF
BrainPET 7T (this system) | 40 cm 15cm ∅ | 24.8 cm | currently 780 ps | 414 kcps @ 12. kBq/ml | 24%
BrainPET 3T1 | 70 cm 20cm ∅ | 19.2 cm | 3 – 6 ns | 31 kcps @ 7.4kBq/ml | 38-42%
NeuroExplorer2 | 70 cm 20cm ∅5 | 49.5 cm | 236 ps | 1.3 Mcps @ 58kBq/ml | 35%
Brain PET insert for 7T (FPGA-only digitization, SNUCM)3,4 | 70 cm 20cm ∅ | 16.7 cm | 1.2 ns | 12 kcps @ 2 kBq/ml | 57%
1Kolb, Armin, et al. European radiology 22 (2012): 1776-1788.
2Li, Hongdi, et al. "Performance Evaluation of a Next-generation Human Brain PET Imager: the NeuroEXPLORER." (Soc Nuclear Med 2023): P801-P801.
3Park, Haewook, et al. "NEMA Performance Evaluation of Long Axial Field-of-View Brain PET Insert with 7T MRI Compatibility." (Soc Nuclear Med 2020): 12-12.
4Won, Jun Yeon, et al. IEEE Transactions on Medical Imaging 40.6 (2021): 1579-1590.
5Private communication
Start activity: ≈ 110 MBq (11C)
Comparison to other systems for neuroscientific applications

## Slide 26: Sensitivity with RF Coil
NEMA: 70 cm long capillary, 1-5 aluminium sleeves
11.1%
3.8%
 | aFOV | ∅ | Scintillator thickness | Energy window | Sensitivity at centre
BrainPET 7T (this system) | 24.8 cm | 39 cm | 2.4 cm | 420-650 keV | 57.2 kcps/MBq
BrainPET 3T | 19.2 cm | 39 cm | 2 cm | 420-650 keV | 20.5 kcps/MBq
NeuroExplorer1 | 49.5 cm | 52.4 cm | 2 cm | unknown | 46.0 kcps/MBq
Brain PET insert for 7T (FPGA-only digitization, SNUCM)2 | 16.7 cm | 33 cm | 2 cm | 450-650 keV | 18.9 kcps/MBq
1Li, Hongdi, et al. "Performance Characteristics of the NeuroEXPLORER, a Next-Generation Human Brain PET/CT Imager." Journal of Nuclear Medicine (2024).
2Won, Jun Yeon, et al. IEEE Transactions on Medical Imaging 40.6 (2021): 1579-1590.
Activity: ≈ 15 MBq (18F),
acquisition times: ≈ 90 seconds
Comparison to other systems for neuroscientific applications
NeuroExplorer1

## Slide 27: Spatial Resolution (Without Corrections!)
No normalization, no random correction, no attenuation correction, no scatter correction
 | Crystal footprint & length (mm3) | Crystals, APDs / SiPMs | DOI | resolution
BrainPET 7T (this system) | 2.02×24 | 196080 17280 | 3-layer 7/8/9 mm | 1.8 mm (1.6 mm?)
BrainPET 3T1 | 2.52×20 | 27648 1728 | No | 2.5 - 5 mm
NeuroExplorer2 | 1.56×3.07×20 | 131328 65664 | light sharing ∽ 3 mm | 1.6 mm
Brain PET insert for 7T (FPGA-only digitization, SNUCM)3 | 2.092×20 | 31225 13824 | 2-layer 8/12 mm | 2.5 - 4 mm
Activity: ≈ 22 MBq
(avg. act. conc. ≈ 40 kBq/ml,18F)
acquisition times: ≈ 2 min.
1Kolb, Armin, et al. European radiology 22 (2012): 1776-1788.
2Li, Hongdi, et al. "Performance Evaluation of a Next-generation Human Brain PET Imager: the NeuroEXPLORER." (Soc Nuclear Med 2023): P801-P801.
3Park, Haewook, et al. "NEMA Performance Evaluation of Long Axial Field-of-View Brain PET Insert with 7T MRI Compatibility."
(Soc Nuclear Med 2020): 12-12.
2.0 mm
1.8 mm
1.6 mm
1.2 mm
7 cm
1.0 mm
1.4 mm
(length: 6.6 cm)
Isocentre
Off-centre
@ ≈ 10 cm from isocentre
Single slice
average
Reconstruction with PRESTO
(ML-EM, 120 iterations, voxel size 0.652 × 1 mm3)
*J. Scheins, et al., IEEE transactions on medical imaging 30.3 (2011): 879-892.
Comparison to other systems for neuroscientific applications

## Slide 28: BrainPET 7T – Ultra High-Field MR Compatible Components
Heat sink with minimized eddy currents
Heat sink for acquisition electronics
Image curtesy by RWTH Aachen & Hyperion Imaging Systems GmbH
Hyperion readout electronics
12×12 dSiPM array with steering electronics
248 mm
RF Blocking feed-troughs
Carbon Fibre Housing
Liquid cooling
850 mm
Scintillation detector
49 mm

## Slide 29: Sensitivity with RF Coil
NEMA: 70 cm long capillary, 1-5 aluminium sleeves
11.1%
3.8%
 | aFOV | ∅ | Scintillator thickness | Energy window | Sensitivity at centre
BrainPET 7T (this system) | 24.8 cm | 39 cm | 2.4 cm | 420-650 keV | 57.2 kcps/MBq
BrainPET 3T | 19.2 cm | 39 cm | 2 cm | 420-650 keV | 20.5 kcps/MBq
NeuroExplorer1 | 49.5 cm | 52.4 cm | 2 cm | unknown | 46.0 kcps/MBq
Brain PET insert for 7T (FPGA-only digitization, SNUCM)2 | 16.7 cm | 33 cm | 2 cm | 450-650 keV | 18.9 kcps/MBq
1Li, Hongdi, et al. "Performance Evaluation of a Next-generation Human Brain PET Imager: the NeuroEXPLORER." (Soc Nuclear Med 2023): P801-P801.
2Won, Jun Yeon, et al. IEEE Transactions on Medical Imaging 40.6 (2021): 1579-1590.
Activity: ≈ 15 MBq (18F),
acquisition times: ≈ 90 seconds
Comparison to other systems for neuroscientific applications
Transmission scan
with 68Ge line source
PET optimized TX/RX coil

## Slide 30: Scintillation Detector with DOI
Bottom Layer
Middle Layer
Top Layer
M. Ito, et al. , 2007 IEEE NSS-MIC Conference Record. Vol. 6. IEEE, 2007.
Lenz, Mirjam. Design and characterisation of an MRI compatible human brain PET insert by means of simulation and experimental studies. PhD Thesis, 2021.
each layer shifted by pixel pitch along x or y direction
Flood map (irradiation of entire detector)

## Slide 31: Positioning Performance (Single Block)
COG positions:
singles full 3x3 neighbourhood, (SiPM)
no energy filter
Signal distributions (avalanche counts):
singles, full 3x3 SiPM neighbourhood, (SiPM)
only photopeak events
script for automatic extraction
SiPM index
Scintillation pixel index
Centre of gravity (CoG) and Maximum Likelihood Positioning, acquisition with 68Ge line source
Scintillation
pixel counts:
all singles
only photopeak events
1st layer (coupled to SiPMs)
2nd layer (middle)
3rd layer (upper)

## Slide 32: Energy Resolution
(and Timing Resolution)
Energy res.: 13% FWHM
Peak position: 509.2 keV
(system wide, measured with 68Ge Line source)
1st layer (coupled to SiPMs)
2nd layer (middle)
3rd layer (upper)
Photopeak positions (avalanches) of single detector block
CRT: 780ps (FWHM)
(system wide, measured with 68Ge line source, skew correction between blocks, but not between SiPMs)
Scintillation crystal ID
Photopeak position [avalanches]

## Slide 33: Aim: Improve Quantitation Accuracy
Target: Reduce quantitation errors caused by data correction to 1%
Boellaard, Ronald. "Standards for PET image acquisition and quantitative data analysis." Journal of nuclear medicine 50.Suppl 1 (2009): 11S-20S.
Quantitation errors of up to 50% (FDG, SUV)
reconstruction bias at low counts (~20%)
Dead time correction method (~7%)
Attenuation based on conventional (polychromatic) CT (~5%)
Single Scatter Correction (~ 15%)
Triple Random/Prompt gamma contamination (~3%)
Not considered in this study:

## Slide 34: PET - Challenge ML-EM reconstruction Bias at low counts
Reilhac, Anthonin, et al. "Simulation-based evaluation of OSEM iterative reconstruction methods in dynamic brain PET studies." Neuroimage 39.1 (2008): 359-368.

## Slide 35: PET - Challenge Me-EM reconstruction Bias
Validation of constant true coincidences count framing scheme
Decay experiment with hot sphere phantom
Hot (test) ROI
Warm (ref) ROI
Hot (test) ROI
Warm (ref) ROI
Activity Concentration
Expected behaviour
T4 = 4th half-life of 11C
(≈ 60 min p.i to ≈ 80 min. p.i.)
Reconstruction frame length
C. Régio Brambilla, et al. "Bias evaluation and reduction in 3D OP-OSEM reconstruction in dynamic equilibrium PET studies with 11C-labeled for binding potential analysis." Plos one 16.1 (2021): e0245580.
Expected behaviour

## Slide 36: Image Reconstruction Bias at low counts
Problem of Iterative Image Reconstruction methods, e.g. Maximum Likelihood Expectation Maximisation
Volunteer 1
Volunteer 2
discard
discard
Task start
baseline
Maintain true coincidence count in reconstruction frames constant
Count rate levels for different subjects in same study ([11C]ABP688 for mGluR5 binding)

## Slide 37: Controlling bias with constant counts
Comparison smokers/non-smokers (healthy control group of SZ study)
With courtesy of C. Regio Brambilla, INM-4, Forschungszentrum Juelich GmbH
Const. 2 min. frames
Const. true coincidence
counts scheme!
constant 11C concentration (decay corrected)
-20 %
-2 %
Régio Brambilla, et al. Plos one 16.1 (2021): e0245580.
F(27,1) = 15.50 and p = 0.001
F(30,1) = 74.72, p < 0.001
Régio Brambilla, et al. Translational Psychiatry 12.1 (2022): 6.

## Slide 38: Improved Dead Time Correction (Block pair-wise)
Issa, Ahlam Said Mohamad, et al. "A detector block-pairwise dead time correction method for improved quantitation with a dedicated BrainPET scanner." Physics in Medicine & Biology67.23 (2022): 235004.

## Slide 39: Positron Range Correction (Chong Li, HCOC-Fellow)
Relevant for non-standard positron emitters and systems with high spatial resolution
9 mm
20 mm
＜2 mm
Maximum range human tissue
Shah, N Jon et al. PloS one, 9(4), 2014
1.6mm

## Slide 40: Positron Range Correction - Approach
Fast raytracing method
Slow down is random process!
Li, Chong, et al. "Fast 3D kernel computation method for positron range correction in PET." Physics in Medicine & Biology 68.2 (2023): 025004.
Account for random process in heterogeneous media with applied magnetic field
511keV
Tissue dependency - Attenuation coefficient µ of 511 keV gamma photon:
Tissue types and location
Linear stopping power:
— velocities of positron
— electron rest energy
— mean excitation energy
— density-effect correction
— velocities of light
— cross section of total (tot), atomic photo-effect (ph), Rayleigh (coh) scattering, Compton (incoh) scattering, and electron-positron production (pair)
— atomic mass unit

## Slide 41: Results – Annihilation probabilities in Test Voxels
120I
18F
1
33
18
120I – water / 18F – Lung
120I – bone / 18F – water
Tissue boundary
Positron emitter
Geant4
Proposed method
Li, Chong, et al. "Fast 3D kernel computation method for positron range correction in PET." Physics in Medicine & Biology 68.2 (2023): 025004.

## Slide 42: Results – Uncertainties
Relevant relative differences of annihilation probabilities 20%
First voxel in water

## Slide 43: CT Attenuation Template
Using newest Photon Counting (PC) CT technology for lowest Bias and potential bone quantification improvement
Siemens PC CT
Electron density phantom for HU to µ conversion
Use virtual monochromatic CT ⇒ avoid scaling from polychromatic average to 511keV
Blood vessels with iodinated blood
Native CT angiography
Virtual non-contrast image
Extra: Material decomposition allows to suppress iodine contrast agent!
⇒ Scans of entire head at high dose (lower bias) possible

## Slide 44: Scatter Correction - Motivation
Ma, Bo, et al. "Scatter correction based on GPU-accelerated Full Monte Carlo simulation for Brain PET/MRI." IEEE transactions on medical imaging 39.1 (2019): 140-151.
Improved Monte Carlo Simulation (MC) Scatter correction
High performance computation required
FET-PET
F18-FET
distribution
simulated
positions of
scatter vertices
preliminary GPU-based simulations running (single GPU):
  throughput/€ ≈ >10 better for GPU vs. CPU
full MC scatter simulation expected time ≈ <120 s./frame
Comparing GPU vs. GATE CPU single thread provides reported acceleration factors of 400-800 (e.g. Ma et.al, 2020, TMI)
PPS approx. 20x faster than GATE (both single thread)
PPS on high-end machine (96 threads) provides acceleration factors of
1600 without multiplexing (2x faster than GPU wo multiplexing)
24000 with multiplexing (30x faster than GPU wo multiplexing)

## Slide 45: MR-PET Compatibility
Electromagnetic issues (actually all solved)
RF Transmit power: kW range ⇨ Requires RF shielding to assure correct working of PET components.
PET Cassette with RF shielding
Eddy currents
B0
Gx, Gy, Gz
Varying B (B0+(Gx,Gy,Gz)) ⇨ Eddy currents ⇨ Distorts gradient fields (spatial encoding!) ⇨ affects advanced applications like MRS.
RF Interference: =
⇨ shielding -60 dB, Tx power 1 kW leads to 1mA • 1V behind shield
⇨ G strength 15mT/100µs, 1cm2 enclosed area leads to 150 mV noise.
RF Rx coil signal: µV without resonance, mV with resonance ⇨ no clocks close to 𝜔0, low noise power supplies.

## Slide 46: Image Reconstruction Bias at low counts
Problem of Iterative Image Reconstruction methods, e.g. Maximum Likelihood Expectation Maximisation
System matrix
(probability)
Coincidence counts in LOR i
Image voxel value j
(decays)
Always positive (due to Poisson nature)
All detected coincidences
Reconstruction Bias at low counts
Activity concentration
Ligand equilibrium
Pharmacologic challenge
Cognitive task

## Slide 47: PET - Task Effect Analysis of [11C]ABP688 PET
Healthy control subjects, bolus-infusion protocol + MNN Task + comparing against Motor Cortex
Task control region
no effects expected
C. Régio Brambilla, et al. "mGluR5 binding changes during a mismatch negativity task in a multimodal protocol with [11C] ABP688 PET/MR-EEG." Translational Psychiatry 12.1 (2022): 1-10.
p-value | Effect size (Kendall’s W) | 
0.001 | 0.255 | -6 ± 2
0.001 | 0.347 | -11 ± 6
0.001 | 0.307 | -7 ± 3
0.193 | 0.035 | -
p-value | Effect size (Kendall’s W) | 
0.001 | 0.255 | -6 ± 2
0.001 | 0.347 | -11 ± 6
0.001 | 0.307 | -7 ± 3
0.193 | 0.035 | -
Notes: The motor cortex was taken as a task-reference region for the analysis because, due to the nature of the MMN elicitation, it is not expected to be activated or show tracer displacements during the MMN task.

## Slide 48: Positioning Method (Maximum Likelihood)
Combined algorithm for determining hit scintillator pixel and calibrated energy
detected charges i.e. triggered SiPM
Scintillating pixel and
scintillator array pixel
: probability for detecting light from scintillator pixel in SiPM
Photodetector pixel (SiPM)
Measured charge distribution
Likelihood values for all crystals
Charge distribution with highest
C. Lerche et al., Physics in Medicine & Biology 61.4 (2016): 1650.

## Slide 49: Current Work (DFG proposal in Evaluation)
Implementation of B0 Field
Compression along the vertical direction
Larger compression for high energy isotopes
18F
68Ga
82Rb
Image with Courtesy of C. Li