## Slide 1
3-Channel SSVEP Brain–Computer Interface with AI-Guided Adaptive Questioning for Non-Verbal Patients
Hardware front-end design · Frame-locked visual stimulation · Filter-bank CCA decoding · LLM-generated dialogue
GGSB conference presentation · Ana Kakniashvili, Anastasia Chikovani · September 2026
Notes: 0:00–0:40. Introduce yourself and the team. One sentence: we built an end-to-end system that lets a paralysed patient answer clinically relevant questions by looking at flickering squares, and an AI turns 3–4 answers into a summary for the doctor. Today: the hardware, the software, and how we validate it.

## Slide 2
Clinical motivation
Locked in and late stage ALS patients retain cognition and vision but lose all motor output
Eye-tracking and switch-based AAC fail as ocular motility declines; EEG-based BCIs need no muscle control.
Steady-state visual evoked potentials (SSVEP): gazing at a stimulus flickering at f Hz produces a narrow-band occipital response at f and its harmonics strong SNR, no user training.
In 7 locked-in patients, all reached ≥70 % accuracy with SSVEP vs. 3 of 7 with P300, at lower workload (Combaz et al., 2013).
Calibration-free asynchronous SSVEP spelling has been shown in a locked-in ALS patient: >90 % accuracy, 22 bits/min (Shi et al., 2020).
Gap: clinical systems cost hundreds to thousands of dollars and require an operator. We target a low cost, self guided device for bedside triage of patient needs.
7 / 7
locked-in patients ≥70 % with SSVEP (Combaz 2013)
>90 %
free-spelling accuracy, locked-in ALS (Shi 2020)
3-Channel SSVEP BCI for Patient Communication · 2
Notes: 0:40–1:40. Establish the need and why SSVEP: it is the paradigm that works best in locked-in cohorts, needs no training, and the response is a clean spectral peak we can detect with very few electrodes. Then state the gap: cost and operator dependence. Our contribution is the integration: cheap hardware + adaptive dialogue.

## Slide 3
System architecture
Closed loop: stimulus → EEG → decoding → adaptive question → clinician summary
Patient gazes at one of 3 flickering squares
8 / 10 / 12 Hz
3-ch occipital EEG
O1 · Oz · O2
analog front end
ADC 250 Hz
2-s epochs
FBCCA decoder
3 sub-bands · 2 harmonics · vote
LLM generates next question + 3 options (JSON)
3–4 iterations
Structured report to clinician:
needs, comfort state, urgency
Software stack: PsychoPy (stimulus, frame-locked) · isolated acquisition process · MATLAB/EEGLAB FBCCA via MATLAB Engine · Python orchestration · Gemini LLM with strict JSON schema
3-Channel SSVEP BCI for Patient Communication · 3
Notes: 1:40–2:30. Walk left to right. Emphasise the loop: unlike a speller, we ask few, high-value questions. Each answer conditions the next question; after 3–4 iterations the system emits a structured summary for the doctor. Mention that decoding and stimulus timing run in separate processes so Bluetooth I/O never drops a frame.

## Slide 4
Hardware: analog front end (per channel)
Instrumentation amplifier → high-pass → anti-alias low-pass → gain → level shift → ADC
AD620
in-amp
Rg = 1 kΩ
G ≈ 50
Sallen-Key HPF
2nd order
1 µF · 160 kΩ
fc ≈ 1 Hz
Sallen-Key LPF
47 kΩ · 117/100 nF
fc ≈ 31 Hz
Sallen-Key LPF
51 kΩ · 150/22 nF
fc ≈ 54 Hz
Gain stage
TL084
G ≈ 40
AC couple + bias
10 µF · 100k/100k
to ADC
Total gain ≈ 2 000 · passband ≈ 1–30 Hz · ±5 V dual supply · three identical channels (U1/U3, U4/U6, U7/U9)
Design rationale
In-amp rejects the common-mode 50 Hz mains voltage on the body (AD620 CMRR ≥ 100 dB) and amplifies the µV differential EEG.
HPF removes electrode DC offset and drift; LPFs suppress EMG and alias energy before sampling.
Gain formula G = 1 + 49.4 kΩ / Rg; Sallen-Key fc = 1/(2πR√(C₁C₂)) — all values verified analytically.
Driven-right-leg (DRL) bias circuit
Common-mode voltage sensed via 470 kΩ network → inverting integrator (1 MΩ ‖ 1 nF) → reference electrode through 390 kΩ.
Negative feedback actively cancels mains pickup; raises effective CMRR by 20–40 dB (Winter & Webster, 1983).
390 kΩ bounds patient auxiliary current to ≈13 µA at 5 V (IEC 60601-1: 10 µA CF / 100 µA BF).
3-Channel SSVEP BCI for Patient Communication · 4
Notes: 2:30–3:40. This is the schematic in words. Say the numbers confidently: gain 50 in the in-amp, 1 Hz high-pass, cascaded low-passes, ×40, level shift. Explain that the DRL is the part most hobby designs skip — it is what makes the signal usable in a 50 Hz environment. Mention safety framing: current-limited, battery-powered research prototype, not a certified device.

## Slide 5
Hardware: design review and revision plan
Quantitative review of the first schematic identified four corrections before fabrication
Parameter | Rev. A (current) | Issue | Rev. B (planned)
First-stage gain | 50 (Rg = 1 kΩ) | ±300 mV electrode offset × 50 exceeds ±5 V rails → saturation | G ≈ 5–10 before HPF; gain ×100–400 after
Low-pass corner | ≈31 Hz (effective) | Removes 2nd/3rd harmonics of 10–12 Hz targets used by FBCCA | fc ≈ 60–80 Hz analog; sharp digital filtering
ADC interface | Bias 2.5 V, ±5 V swing | 3.3 V ADC mid-scale is 1.65 V; over-voltage risk | 3.3 V rail-to-rail buffer at 1.65 V + clamps
Op-amp noise | TL084 (≈18–37 nV/√Hz) | Noise comparable to SSVEP amplitude (<1 µV) | OPA4192 / OPA1644 (≈5 nV/√Hz)
Input protection | None | Patient safety and ESD | 10–47 kΩ series R + clamps; battery + BLE isolation
Method: each stage recomputed analytically (gain, corner frequency, headroom); compared against AD620/AD8232 application notes and the OpenEEG ModularEEG reference design (in-amp G = 12 before HPF, total G ≈ 7 800). Corrections are being validated in simulation prior to Rev. B layout.
3-Channel SSVEP BCI for Patient Communication · 5
Notes: 3:40–4:40. This slide earns credibility: you found your own problems. Reviewers will ask about saturation, noise and safety anyway — pre-empt them. Key sentence: 'The topology is right; the gain distribution and passband were wrong for SSVEP, and we corrected both before fabrication.' Do not oversell Rev. B as done if it is not.

## Slide 6
Software I: visual stimulation and acquisition
Frame-locked flicker in PsychoPy with hardware I/O isolated in a separate process
A · 8 Hz
"I have called the nurse.
What else can I do?"
B · 10 Hz
C · 12 Hz
Screen layout: three 300 px squares at wide angular separation; AI prompt centred; option labels beneath each square.
Flicker is generated per frame: square on when (frame mod P) < P/2, P = round(f_refresh / f). Exact integer periods (60 Hz → 6 and 5 frames for 10 and 12 Hz) give phase-stable square-wave stimulation.
Trial length 5 s per question; frequencies 8 / 10 / 12 Hz chosen ≥2 Hz apart, no mutual low-order harmonics, none coincident with 50 Hz sub-harmonics.
Acquisition runs in an isolated worker process (multiprocessing) so Bluetooth latency cannot cause dropped frames in the stimulus loop.
Current validation platform: g.tec Unicorn, 8 ch, 250 Hz; occipital PO7 / Oz / PO8 used → identical channel count to the custom board. Data streamed as float32, converted to BDF (MNE) for analysis.
3-Channel SSVEP BCI for Patient Communication · 6
Notes: 4:40–5:40. Explain frame-locking: the display refresh is the clock, so each frequency is an integer number of frames. Note honestly that 8 Hz on a 60 Hz display rounds to 8 frames = 7.5 Hz actual, so Rev. B will move to 7.5 or 8.57 Hz or verify with a photodiode. Explain the two-process architecture — it is a real engineering decision. State that the decoding pipeline is validated on a commercial headset now, and the custom board will be substituted once its noise floor and CMRR are measured.

## Slide 7
Software II: filter-bank CCA decoding
Training-free target identification exploiting fundamental and harmonic responses (Chen et al., 2015)
Pre-processing
discard warm-up · 5–40 Hz
4th-order Butterworth
Epoching
2-s windows
3 channels
Filter bank
5–15 · 15–25 · 25–40 Hz
CCA per band vs.
sin/cos refs, h = 1,2
for 8, 10, 12 Hz
Weighted sum
ρ = Σ w_k ρ_k
argmax → vote
Formulation
Reference set for target f: Y_f = [sin(2πft), cos(2πft), sin(4πft), cos(4πft)]ᵀ
Sub-band k: ρ_k(f) = max corr( Wₓᵀ X_k , W_yᵀ Y_f )
Fusion: ρ(f) = Σ_k w_k · ρ_k(f), w_k = k⁻¹ (k = 1…3)
Decision per epoch: f̂ = argmax_f ρ(f); trial decision = majority vote over epochs
Spatial filters Wₓ are learned per epoch from the data — no calibration session, robust to inter-subject variability.
Why filter-bank CCA
Harmonic sub-bands add discriminative SNR; FBCCA outperforms plain CCA and reached 151 bits/min online in a 40-target speller (Chen 2015).
Three targets and 2-s epochs give a low-complexity classifier suited to bedside use and to on-device execution in future revisions.
Vote count per target [n₈, n₁₀, n₁₂] is returned to Python; margin between top two counts serves as a confidence measure for an idle / 'repeat question' state.
Implementation: MATLAB + EEGLAB (canoncorr), called from Python via MATLAB Engine; Welch PSD inspected per trial.
3-Channel SSVEP BCI for Patient Communication · 7
Notes: 5:40–6:50. Give the algorithm its due: CCA finds the spatial combination of the three channels that best correlates with a sine/cosine template at each candidate frequency and its second harmonic; doing this in three sub-bands and weighting them is FBCCA. No training. The output per trial is a vote count; the winning option is fed to the LLM. Mention the confidence margin as the hook for asynchronous operation.

## Slide 8
Software III: LLM-guided adaptive dialogue
Each decoded answer conditions the next question the loop terminates when needs are resolved
Context prompt
"What is the patient's primary complaint?"
LLM (Gemini)
strict JSON: display_text,
3 options, is_finished
Render options as
flicker labels;
5-s SSVEP trial
FBCCA → chosen option
appended to context
is_finished → summary
for clinician
Example session (3 iterations)
Q1 "What do you need?" → Pain · Water · Nurse
A1 Pain (12 Hz, 2 of 2 epochs)
Q2 "I have alerted the nurse. Where is the pain?" → Head · Chest · Back
A2 Chest (10 Hz)
Q3 "Is it hard to breathe?" → Yes · No · Unsure
A3 Yes (8 Hz) → is_finished = true
Report: chest pain with dyspnoea; urgent clinical review recommended.
Design constraints
Structured output enforced (response_mime_type = application/json) and validated; labels truncated to ≤20 characters, display text ≤12 words, exactly 3 options.
Retry logic and a fallback question guarantee the interface never blocks on a model error.
Options are mapped to fixed frequencies A/B/C so the patient learns a stable spatial layout.
Final summary is decision support only: it lists reported symptoms and the answer path; it does not diagnose.
3-Channel SSVEP BCI for Patient Communication · 8
Notes: 6:50–7:50. This is the novel part of the system — most SSVEP BCIs are spellers. Explain that the LLM's job is to compress the conversation: three well-chosen questions replace thirty spelled characters. Stress the guard-rails: strict schema, bounded label length, no diagnostic claims. Walk through the example quickly.

## Slide 9
Validation protocol and metrics
Bench characterisation of the front end, then physiological and closed-loop evaluation
1 · Bench
Input-referred noise (shorted inputs), µVrms, 0.5–100 Hz
Gain and −3 dB corners vs. design (signal generator + attenuator)
CMRR at 50 Hz with DRL on / off
2 · Physiological
Eyes-closed alpha blocking at Oz (8–13 Hz)
Welch PSD peak at each stimulus frequency and 2nd harmonic
SNR of SSVEP vs. neighbouring bins
3 · Closed loop
Offline accuracy and ITR (bits/min) across subjects and epoch lengths
Confusion matrix incl. idle state
End-to-end task success: correct 3-question path
Reference points from the literature
Low-cost 2-channel ESP32 SSVEP BCI: 95.6 ± 3.7 % accuracy, 102 bits/min (Teversham et al., 2022) · ADS1298-class front ends: ≈0.5 µVrms input noise · target for this device: <2 µVrms, >85 % 3-class accuracy at 2 s, ITR reported with 95 % CI.
3-Channel SSVEP BCI for Patient Communication · 9
Notes: 7:50–8:40. If you have measured numbers, put them here (noise floor, a PSD figure, an accuracy table). If not yet, present this as the protocol and be explicit about which stages are complete. Reviewers respect a clear plan more than a vague claim. Name the benchmarks you will compare against.

## Slide 10
Limitations and next steps
Known constraints of Rev. A and the features planned to differentiate Rev. B
Current limitations
12-bit MCU ADC limits resolution (~0.4 µV/LSB at G = 2 000); 24-bit ΔΣ (ADS131M04) planned.
Display-locked stimulation quantises frequencies to divisors of the refresh rate (8 Hz → 7.5 Hz at 60 Hz).
Decoding depends on a MATLAB Engine bridge; on-device or Python-native FBCCA needed for portability.
Evaluation so far on healthy volunteers with a commercial headset; no patient data yet.
Research prototype: not IEC 60601-1 certified; battery operation and isolation required for any human use.
Planned innovations
Onboard µV self-test signal: device verifies its own gain and noise floor at start-up.
Electrode-contact quality readout via low-current AC injection and lock-in demodulation.
Photodiode stimulus-sync channel: true flicker phase, frame-drop rejection, phase-coded targets.
EOG blink channel as volitional confirm / cancel to reduce false selections.
Confidence-thresholded asynchronous decoding: explicit 'no answer' state drives re-prompting.
3-Channel SSVEP BCI for Patient Communication · 10
Notes: 8:40–9:30. Owning limitations increases credibility. Then pivot to what makes Rev. B novel: a self-checking, stimulus-synchronised, asynchronous device rather than a bare amplifier. Keep each item to one sentence.

## Slide 11
Conclusion
An end-to-end SSVEP communication system: 3-channel analog front end with DRL, frame-locked stimulation, training-free FBCCA decoding, and LLM-generated adaptive questioning with a structured clinician report.
Quantitative self-review corrected gain distribution, passband and ADC interfacing before fabrication.
Positioned between hobby boards and clinical systems: inspectable discrete design, sub-$50 BOM target, few high-value questions instead of spelling.
Key references
Combaz A. et al. (2013) PLoS ONE 8(9): e73691 · Shi N. et al. (2020) J. Neurorestoratology 8(1): 40–52
Lin Z. et al. (2006) IEEE TBME 53(12): 2610–2614 · Chen X. et al. (2015) J. Neural Eng. 12(4): 046008
Teversham J. et al. (2022) IEEE EMBC, 208–213 · Winter B.B. & Webster J.G. (1983) IEEE TBME 30(1): 62–66
Okahara Y. et al. (2018) Sci. Rep. 8: 16787 · OpenEEG ModularEEG design documentation
Thank you — questions welcome
Notes: 9:30–10:00. Three sentences, then stop. Anticipated questions: (1) noise floor and CMRR numbers; (2) how you avoid LLM hallucination — strict schema, no diagnosis, human clinician in the loop; (3) why not ADS1299 — cost, pedagogy, inspectability, and we treat it as the benchmark; (4) safety — battery, isolation, current limiting, not certified.